GDM-8245 Specifications

The specifications apply when the GDM-8245 is powered on for at least 30 minutes under +20°C~+30°C.

| DC VOLTAGE |
| --- |
| Range | 500mV, 5V, 50V, 500V, 1000V 5 ranges |
| Accuracy | +(0.03% rdg + 4 digits) |
| Input Impedance | 10MΩ |
| AC VOLTAGE True RMS (AC or AC+DC True RMS) |
| Range | 500mV, 5V, 50V, 500V, 1000V 5 ranges |
| Accuracy | 500mV ~ 50V 3 ranges: 20Hz ~ 45Hz: +(1% rdg + 15 digits) 45Hz ~ 2kHz: +(0.5% rdg + 15 digits) 2kHz ~ 10kHz: +(1% rdg + 15 digits) 10kHz ~ 20kHz: +(2% rdg + 30 digits) 20kHz ~ 50kHz: +(5% rdg + 30 digits) 500V,1000V 2 ranges: 45Hz ~ 1kHz: +(0.5% rdg + 15 digits) |
| Input Impedance | 10MΩ |
| DC CURRENT |
| Range | 500μA, 5mA, 50mA, 500mA, 2A, 20A 6 ranges |
| Accuracy | 500μA ~ 500mA 4 ranges：+(0.02% rdg + 2 digits) |
|  | 2A ~ 20A 2 ranges：+(0.3% rdg + 2 digits) |
| AC CURRENT True RMS (AC or AC+DC True RMS) |
| Range | 500μA, 5mA, 50mA, 500mA, 2A, 20A 6 ranges |
| Accuracy | 500μA ~ 20A 6 ranges 20Hz ~ 45Hz: +(1% rdg + 15 digits) 45Hz ~ 2kHz: +(0.5% rdg + 15 digits) 500μA ~ 50mA 3 ranges 2kHz ~ 10kHz: +(1% rdg + 15 digits) 10kHz ~ 20kHz: +(2% rdg + 15 digits) |
| RESISTANCE |
| Range | 500Ω, 5kΩ, 50kΩ, 500kΩ, 5MΩ, 20MΩ 6 ranges |
| Accuracy | 500Ω range：+(0.1% rdg + 4 digits) 5kΩ~ 500kΩ 3 ranges：+(0.1% rdg + 2 digits) 5MΩ range：+(0.2% rdg + 2 digits) 20MΩ range：+(0.3% rdg + 2 digits) |
| DIODE TEST |
|  | Maximum forward voltage 1.5V, open voltage 2.8V |
| CAPACITANCE |
| Range | 5nF, 50nF, 500nF, 5μF, 50μF 5 ranges |
| Accuracy | +(2% rdg + 4 digits) |
| FREQUENCY |
| Input Level (Sine Wave) | mV range： 10Hz ~ 50kHz >120mV 50kHz ~ 150kHz >200mV 5V ~ 50V range： 10Hz ~ 200kHz >1.2V 500V range： 20Hz ~ 1kHz >12V |
| FUNCTIONS |  |
|  | Auto Range, Manual Range, Max, Min, dBm, Rel, Hold |
| CONTINUITY BEEP |  |
|  | < 5Ω |
| DISPLAY |  |
|  | Dual display 0.4", 0.5" 7 segment LED |
| POWER SOURCE |  |
|  | AC 100V/120V/230V +10%, 50/60Hz |
| DIMENSIONS & WEIGHT |  |
|  | 251(W) × 91(H) × 291(D) mm Approx 2.6kg |