GDM-8255A Specifications

The specifications apply when the GDM-8255A is powered on for at least 30 minutes under +20°C~+30°C.

|  | GDM-8255A |
| --- | --- |
| FULL SCALE |
|  | 199,999 Counts |
| SAMPLAE RATE |
| Slow | 5 1/2 digits, 10 readings/second |
| Medium | 4 1/2 digits, 30 readings/second |
| Fast | 3 1/2 digits, 60 readings/second |
| DC VOLTAGE |
| Range | 100mV, 1V, 10V, 100V, 1000V 5 ranges |
| Accuracy | ±(0.012% rdg + 5 digits) |
| Input Impedance | 10MΩ |
| AC VOLTAGE True RMS |
| Range | 100mV, 1V, 10V, 100V, 750V 5 ranges |
| Accuracy | 20Hz~45Hz : ±(1% rdg + 100 digits) 45Hz ~ 10kHz : ±(0.2% rdg + 100 digits) 10kHz ~ 30kHz :±(1% rdg + 100 digits) 30kHz~100kHz : ±(3% rdg + 200 digits) |
| Input Impedance | 1.1MΩ in parallel with approx. 100pF |
| DC CURRENT |
| Range | 10mA, 100mA, 1A, 10A 4ranges |
| Accuracy | 10mA , 100mA : ±(0.05% rdg + 15 digits) 1A, 10A range : ±(0.2% rdg + 5 digits) |
| AC CURRENT True RMS |
| Range | 10mA, 100mA, 1A, 10A 4ranges |
| Accuracy | 10mA , 100mA range: 20Hz ~ 50Hz : ±(1.5% rdg + 100 digits) 50Hz ~ 10kHz : ±(0.5% rdg + 100 digits) 10kHz~20kHz : ±(2% rdg + 200 digits) 1A , 10A range: 50Hz~10kHz : ±(1% rdg + 100 digits) |
| RESISTANCE |
| Range | 100Ω, 1kΩ, 10kΩ, 100kΩ, 1MΩ, 10MΩ, 100MΩ 7 ranges |
| 2W Accuracy | 100Ω range : ±(0.1% rdg + 8 digits)(*) 1kΩ range : ±(0.08% rdg + 5 digits)(*) 10kΩ ~ 1MΩ 3 ranges : ±(0.06% rdg + 5 digits) 10MΩ range : ± (0.3% rdg + 5 digits) 100MΩ range : ±(3% rdg + 8 digits) |
| 4W Accuracy | 100Ω range : ±(0.05% rdg + 8 digits) 1kΩ ~ 1MΩ 4 ranges : ±(0.05% rdg + 5 digits) 10MΩ range : ±(0.3% rdg + 5 digits) 100MΩ range : ±(3% rdg + 8 digits) |
| DIODE TEST |
|  | Open Circuit Voltage 2.0V Test Current 0.5mA |
| FREQUENCY |
|  | 10Hz~800kHz |
| CONTINUITY BEEPER |
|  | 1 ~ 1000Ω Define by user |
| TEMPERATURE |
|  | 0° C ~ 300° C : J , K , T Type |
| OTHER FUNCTIONS |
|  | Auto range / Manual, Math : MX + B/ % / 1/X Max, Min, dBm, dB, REL, Hold, Compare, Store, Recall |
| DISPLAY |
|  | VFD, Two Colors Display |
| INTERFACE |
|  | Digital I/O, USB Device, RS-232C |
| POWER SOURCE |
|  | AC 100V-240V ±10%, 50/60Hz |
| OPTIONAL |
|  | GDM-SC1 : Multi-channels (16+2) Scanner Card (Up to 2 sets of GDM-SC1 can be installed at same time for one GDM-8255A) |
| DIMENSIONS & WEIGHT |
|  | 265(W) x 107(H) x 350(D) mm Approx. 2.6 kg |
| Note : (*)USE "REL" Mode |