GDM-8342 / GDM-8341 Specifications

The specifications apply when the GDM-8342 / GDM-8341 is powered on for at least 30 minutes under +20°C~+30°C.

| Function | Range (*2) | Resolution | Input Resistance etc. | Accuracy(*1) 1 Year 23°C±5°C |
| --- | --- | --- | --- | --- |
| DC Voltage | 500.00 mV | 10μV | 10MΩ or >10GΩ | 0.02 + 4 |
|  | 5.0000 V | 100μV | 10MΩ or >10GΩ | 0.02 + 4 |
|  | 50.000 V | 1mV | 11.11MΩ | 0.02 + 4 |
|  | 500.00 V | 10mV | 10.1MΩ | 0.02 + 4 |
|  | 1000.0 V | 100mV | 10MΩ | 0.02 + 4 |
| Resistance | 500.00 Ω | 10mΩ | 0.83mA | 0.10 + 5 (*3) |
|  | 5.0000 kΩ | 100mΩ | 0.83mA | 0.10 + 3 (*3) |
|  | 50.000 kΩ | 1Ω | 83μA | 0.10 + 3 |
|  | 500.00 kΩ | 10Ω | 8.3μA | 0.10 + 3 |
|  | 5.0000 MΩ | 100Ω | 830nA | 0.10 + 3 |
|  | 50.000 MΩ | 1kΩ | 560nA//10 MΩ | 0.30 + 3 |
| DC Current | 500.00 μA | 10nA | 0.06V max. | 0.05 + 5 |
|  | 5.0000 mA | 100nA | 0.6V max. | 0.05 + 4 |
|  | 50.000 mA | 1μA | 0.14V max. | 0.05 + 4 |
|  | 500.00 m A | 10μA | 1.4V max. | 0.10 + 4 |
|  | 5.0000 A | 100μA | 0.5V max. | 0.25 + 5 |
|  | 10.000 A | 1mA | 0.8V max. | 0.25 + 5 |
| Continuity | 5000.0 Ω | 100mΩ | 0.83mA | 0.10 + 5 |
| Diode Test | 5.0000 V | 100μV | 0.83mA | 0.05 + 5 |
| Capacitance | 5.000 nF: 0.5~1nF | 0.001nF | 8.3μA | 2.00 + 20 |
|  | 5.000 nF: 1~5nF | 0.001nF | 8.3μA | 2.00 + 10 |
|  | 50.00 nF: 5~10nF | 0.01nF | 8.3μA | 2.00 + 30 |
|  | 50.00 nF: 10~50nF | 0.01nF | 8.3μA | 2.00 + 10 |
|  | 500.0 nF | 0.1nF | 83μA | 2.00 + 4 |
|  | 5.000 μF | 1nF | 0.56mA | 2.00 + 4 |
|  | 50.00 μF | 10nF | 0.83mA | 2.00 + 4 |
| True RMS Volatge AC (or AC+DC) (*4/*5/*6) | 500.00 mV | 10μV | 30Hz – 50Hz | 1.00 + 40 |
|  |  |  | 50Hz – 10kHz | 0.50 + 40 |
|  |  |  | 10kHz – 30kHz | 2.00 + 60 |
|  |  |  | 30kHz – 100kHz | 3.00 + 120 |
|  | 5.0000 V | 100μV | 30Hz – 50Hz | 1.00 + 20 |
|  |  |  | 50Hz – 10kHz | 0.35 + 15 |
|  |  |  | 10kHz – 30kHz | 1.00 + 20 |
|  |  |  | 30kHz – 100kHz | 3.00 + 50 |
|  | 50.000 V | 1mV | 30Hz – 50Hz | 1.00 + 20 |
|  |  |  | 50Hz – 10kHz | 0.35 + 15 |
|  |  |  | 10kHz – 30kHz | 1.00 + 20 |
|  |  |  | 30kHz – 100kHz | 3.00 + 50 |
|  | 500.00 V | 10mV | 30Hz – 50Hz | ------ |
|  |  |  | 50Hz – 10kHz | 0.50 + 15 |
|  |  |  | 10kHz – 30kHz | 1.00 + 20 |
|  |  |  | 30kHz – 100kHz | 3.00 + 50 |
|  | 750.0 V | 100mV | 30Hz – 50Hz | ------ |
|  |  |  | 50Hz – 10kHz | 0.50 + 15 |
|  |  |  | 10kHz – 30kHz | ------ |
|  |  |  | 30kHz – 100kHz | ------ |
| True RMS Current AC (or AC+DC) (*4/*5/*6) | 500.00 μA | 10nA | 30Hz – 50Hz | 1.50 + 50 |
|  |  |  | 50Hz – 2kHz | 0.50 + 40 |
|  |  |  | 2kHz – 5kHz | 1.50 + 50 |
|  |  |  | 5kHz – 20kHz | 3.00 + 75 |
|  | 5.0000 mA | 100nA | 30Hz – 50Hz | 1.50 + 40 |
|  |  |  | 50Hz – 2kHz | 0.50 + 20 |
|  |  |  | 2kHz – 5kHz | 1.50 + 40 |
|  |  |  | 5kHz – 20kHz | 3.00 + 60 |
|  | 50.000 mA | 1μA | 30Hz – 50Hz | 1.50 + 40 |
|  |  |  | 50Hz – 2kHz | 0.50 + 20 |
|  |  |  | 2kHz – 5kHz | 1.50 + 40 |
|  |  |  | 5kHz – 20kHz | 3.00 + 60 |
|  | 500.00 mA | 10μA | 30Hz – 50Hz | 1.50 + 40 |
|  |  |  | 50Hz – 2kHz | 0.50 + 20 |
|  |  |  | 2kHz – 5kHz | 1.50 + 40 |
|  |  |  | 5kHz – 20kHz | 3.00 + 60 |
|  | 5.0000 A | 100μA | 30Hz – 50Hz | 2.00 + 40 |
|  |  |  | 50Hz – 2kHz | 0.50 + 30 |
|  | 10.000 A | 1mA | 30Hz – 50Hz | 2.00 + 40 |
|  |  |  | 50Hz – 2kHz | 0.50 + 30 |
| Frequency | 10Hz – 500Hz | ------ | ------ | 0.01 + 5 |
|  | 500Hz – 500kHz |  |  | 0.01 + 3 |
|  | 500kHz – 1MHz |  |  | 0.01 + 5 |
| Temperature (*7) (Thermocouple) | -200 °C ~ +300 °C | 0.1 °C | J / T / K | 2 °C (typical) |
| Display | VFD, Tow Colors Display |
| Interface | USB Device, USB Host (for GDM-8342 only), GPIB (optional for GDM-8342 only) |
| Power Source | AC 100 V / 120 V / 220 V / 240 V ±10%, 50-60Hz |
| Power Consumption | Max. 15VA |
| Dimensions | 265(W) x 107(H) x 302(D) mm |
| Weight | Approx. 2.9kg without option |

1. Accuracy = +/- (% of reading + digits)

2. 2% overrange on all ranges, except 10A is 20% overrange

3. REL function is on

4. The AC+DC is in AC coupled

5. The accuracy of AC+DC is equal to AC with 10 more digits added

6. AC characteristics are for sinewave input > 5% of range

7. Specifications do not include probe accuracy