GDM-8351 Specifications

The specifications apply when the GDM-8351 is powered on for at least 30 minutes under +20°C~+30°C.

| Specification Function(*1) |
| --- |
| Range (*2) | Resolution | Input Resistance etc. | Accuracy(*3) 1 Year (23°C±5°C) |
| DC Voltage |
| 100.000 mV | 1μV | 10MΩ or >10GΩ | 0.012 + 8 |
| 1.00000 V | 10μV | 10MΩ or >10GΩ | 0.012 + 5 |
| 10.0000 V | 100μV | 11.1MΩ | 0.012 + 5 |
| 100.000 V | 1mV | 10.1MΩ | 0.012 + 5 |
| 1000.00 V | 10mV | 10MΩ | 0.012 + 5 |
| Resistance |
| 100.000 Ω | 1mΩ | 1mA | 0.05 + 8 |
| 1.00000 kΩ | 10mΩ | 1mA | 0.05 + 5 |
| 10.0000 kΩ | 100mΩ | 100μA | 0.05 + 5 |
| 100.000 kΩ | 1Ω | 10μA | 0.05 + 5 |
| 1.00000 MΩ | 10Ω | 1μA | 0.05 + 5 |
| 10.0000 MΩ | 100Ω | 0.5μA | 0.30 + 5 |
| 100.000 MΩ | 1kΩ | 0.5μA //10 MΩ | 3.00 + 8 |
| DC Current |
| 10.0000 mA | 100nA | 1Ω | 0.05 + 15 |
| 100.000 mA | 1μA | 1Ω | 0.05 + 5 |
| 1.00000 A | 10μA | 0.1Ω | 0.20 + 5 |
| 10.0000 A | 100μA | 0.01Ω | 0.20 + 5 |
| Continuity |
| 1000.00 Ω | 10mΩ | 1mA | 0.05 + 5 |
| Diode Test |
| 6.0000 V | 100μV | 1mA@6V | 0.05 + 15 |
| Capacitance |
| 10.00 nF | 0.01nF | 10μA | 2.0 + 10 |
| 100.0 nF | 0.1nF | 10μA | 2.0 + 4 |
| 1.000 μF | 0.001μF | 100μA | 2.0 + 4 |
| 10.00 μF | 0.01μF | 1mA | 2.0 + 4 |
| 100.0 μF | 0.1μF | 1mA | 2.0 + 4 |
| True RMS AC (or AC+DC – AC Coupled) Voltage |
| 100.000 mV | 1μV | 20Hz – 45Hz | 1.0 + 100 |
|  |  | 45Hz – 10kHz | 0.3 + 100 |
|  |  | 10kHz – 30kHz | 1.5 + 300 |
|  |  | 30kHz – 100kHz | 5.0 + 300 |
| 1.00000 V | 10μV | 20Hz – 45Hz | 1.0 + 100 |
|  |  | 45Hz – 10kHz | 0.2 + 100 |
|  |  | 10kHz – 30kHz | 1.0 + 100 |
|  |  | 30kHz – 100kHz | 3.0 + 200 |
| 10.0000 V | 100μV | 20Hz – 45Hz | 1.0 + 100 |
|  |  | 45Hz – 10kHz | 0.2 + 100 |
|  |  | 10kHz – 30kHz | 1.0 + 100 |
|  |  | 30kHz – 100kHz | 3.0 + 200 |
| 100.000 V | 1mV | 20Hz – 45Hz | 1.0 + 100 |
|  |  | 45Hz – 10kHz | 0.2 + 100 |
|  |  | 10kHz – 30kHz | 1.0 + 100 |
|  |  | 30kHz – 100kHz | 3.0 + 200 |
| 750.00 V | 10mV | 20Hz – 45Hz | 1.0 + 100 |
|  |  | 45Hz – 10kHz | 0.2 + 100 |
|  |  | 10kHz – 30kHz | 1.0 + 100 |
|  |  | 30kHz – 100kHz | 3.0 + 200 |
| True RMS AC (or AC+DC – AC Coupled) Current |
| 10.0000 mA | 100nA | 20Hz – 45Hz | 1.5 + 100 |
|  |  | 45Hz – 2kHz | 0.5 + 100 |
|  |  | 2kHz – 10kHz | 2.0 + 200 |
| 100.000 mA | 1μA | 20Hz – 45Hz | 1.5 + 100 |
|  |  | 45Hz – 2kHz | 0.5 + 100 |
|  |  | 2kHz – 10kHz | 2.0 + 200 |
| 1.00000 A | 10μA | 20Hz – 45Hz | 1.5 + 100 |
|  |  | 45Hz – 2kHz | 0.5 + 100 |
|  |  | 2kHz – 10kHz | 2.0 + 200 |
| 10.0000 A | 100μA | 20Hz – 45Hz | 1.5 + 100 |
|  |  | 45Hz – 2kHz | 1.0 + 100 |
|  |  | 2kHz – 10kHz | ----- |
| Frequency |
| (Voltage)10Hz – 1MHz | ------ | ------ | 0.01 + 3 |
| (Current)20Hz – 10kHz |  |  |  |
| Temperature (Thermocouple) |
| -200 °C ~ 0 °C | 0.01 °C | J / T / K | 0.4 °C (typical) |
| 0 °C ~ +300 °C |  |  | 0.2 °C (typical) |
| Display |
| VFD, Tow Colors Display |
| Interface |
| RS-232C, USB device (USBCDC & USBTMC) |
| Power Source |
| AC 100 V / 120 V / 220 V / 240 V ±10%, 50-60Hz ; Power Consumption Max. 15VA |
| Dimensions & Weight |
| 265(W) x 107(H) x 302(D) mm ; Approx. 2.9kg |

1. All specifications are applicable to the main (1st) display only and warmed up for at least 30 minutes and operated in the slow rate.
2. 20% overrange on all ranges, except 750V/10A range
3. Accuracy: ± (% of Reading + Digits)