GPT-9800 Specifications

The specifications apply when the GPT-9800 is powered on for at least 30 minutes under +20°C~+30°C.

| Model \ Func. | AC WITHSTANDING | DC WITHSTANDING | INSULATION RESISITANCE | GROUND BOND |
| --- | --- | --- | --- | --- |
| GPT-9804 | ( | ( | ( | ( |
| GPT-9803 | ( | ( | ( | ( |
| GPT-9802 | ( | ( | ( | ( |
| GPT-9801 | ( | ( | ( | ( |

| AC WITHSTANDING |
| --- |
| Output-Voltage Range | 0.050kV~ 5.000kV ac |
| Output-Voltage Resolution | 2V/step |
| Output-Voltage Accuracy | ± (1% of setting +5V) [no load] |
| Maximum Rated Load | 200 VA (5kV/40mA) |
| Maximum Rated Current | 40mA (0.5kV < V ≦ 5kV) 10mA (0.05kV ≦ V≦ 0.5kV) |
| Output-Voltage Waveform | Sine wave |
| Output-Voltage Frequency | 50 Hz / 60 Hz selectable |
| Voltage regulation | ± (1% of reading +5V) [full load → no load] |
| Voltmeter Accuracy | ± (1% of reading+ 5V) |
| Current Measurement Range | 0.001mA~040.0mA |
| Current Best Resolution | 0.001mA / 0.01mA / 0.1mA |
| AC Current Measurement Accuracy | ±(1.5% of reading +30 counts) when HISET < 1.11mA ±(1.5% of reading +3 counts) when HISET ≧ 1.11mA |
| Window Comparator Method | Yes |
| ARC Detect | Yes |
| RAMP (Ramp-Up Time) | 0.1~999.9s |
| TIMER (Test Time)* | OFF, 0.5s~999.9s |
| GND | ON/OFF |
| DC WITHSTANDING |
| Output-Voltage Range | 0.050kV~ 6.000kV dc |
| Output-Voltage Resolution | 2V/step |
| Output-Voltage Accuracy | ± (1% of setting +5V) [no load] |
| Maximum Rated Load | 50W (5kV/10mA) |
| Maximum Rated Current | 10mA (0.5kV < V ≦ 6kV) 2mA (0.05kV ≦ V≦ 0.5kV) |
| Voltage regulation | ± (1% of reading +5V) [full load → no load] |
| Voltmeter Accuracy | ± (1% of reading+ 5V) |
| Current Measurement Range | 0.001mA~010.0mA |
| Current Best Resolution | 0.001mA / 0.01mA / 0.1mA |
| DC Current Measurement Accuracy | ±(1.5% of reading +30 counts) when HISET < 1.11mA ±(1.5% of reading +3 counts) when HISET ≧ 1.11mA |
| Window Comparator Method | Yes |
| ARC Detect | Yes |
| RAMP (Ramp-Up Time) | 0.1~999.9s |
| TIMER (Test Time)* | OFF, 0.5s~999.9s |
| GND | ON/OFF |
| INSULATION RESISITANCE |
| Output-Voltage | 50V~1000V dc |
| Output-Voltage Resolution | 50V/step |
| Output-Voltage Accuracy | ± (1% of setting +5V) [no load] |
| Resistance Measurement Range | 1 Mohm~ 9500 Mohm |
|  | Test voltage | Measurable Range | Accuracy |
|  | 50V≦V≦450V | 1 Mohm~ 50 Mohm 51 Mohm~ 2000 Mohm | ±(5% of reading +1 count) ±(10% of reading +1 counts) |
|  | 500V≦V≦1000V | 1M ohm~ 500 Mohm 501 Mohm~ 9500 Mohm | ±(5% of reading +1 count) ±(10% of reading +1 count) |
| Window Comparator Method | Yes |
| Output Impedance | 600 kohm |
| RAMP (Ramp-Up Time) | 0.1~999.9s |
| TIMER (Test Time) | 1s~999.9s |
| GND | OFF (fixed) |
| GROUND BOND |
| Output-Current Range | 03.00A~30.00A ac |
| Output-Current Resolution | 0.01A |
| Output-Current Accuracy | 3A≦I≦8A : ± (1% of setting +0.2A) 8A<I≦30A : ± (1% of setting +0.05A) |
| Test-Voltage | 6Vac max. (open circuit) |
| Test-Voltage Frequency | 50 Hz / 60 Hz selectable |
| Resistance Measurement Range | 10mW~650.0mW |
| Resistance Measurement Resolution | 0.1mW |
| Resistance Measurement Accuracy | ± (1% of reading +2mW) |
| Window Comparator Method | Yes |
| TIMER (Test Time) | 0.5s~999.9s |
| GND | OFF (fixed) |
| Test Method | Four Terminal |
| MEMORY |
| Single Step Memory | MANU : 100 blocks |
| Automatic Testing Memory | AUTO : 100 blocks, menu per auto : 16 |
| INTERFACE |
| RS232 | Standard |
| USB (Device) | Standard |
| GPIB | Option |
| Remote Terminal (Front) | Standard |
| SIGNAL IO | Standard |
| General |
| DISPLY | 240 x 64 dot matrix LED back light LCD |
| POWER SOURCE | AC100V/120V/220V/230V ±10%, 50/60Hz |
| POWER CONSUMPTION | 500VA max. |
| DIMENSIONS & WEIGHT | 330(W) x 148(H) x 452(D) mm, approx.19kg(Max) |

* The “TIMER OFF” can only be performed when the tester is in the special MANU mode