GSP-9300B Specifications

The specifications apply when the GSP-9300B is powered on for at least 30 minutes under +20°C~+30°C.

| Frequency |
| --- |
| Frequency | Range | 9 kHz to 3 GHz |
|  | Resolution | 1 Hz |
| Frequency Reference | Accuracy | ±(period since last adjustment X aging rate) + stability over temperature + supply voltage stability |
|  | Aging Rate | ±1 ppm max. | 1 year after last adjustment |
|  | Frequency Stability over Temperature | ±0.025 ppm | 0 to 50 °C |
|  | Supply Voltage Stability | ±0.02 ppm |  |
| Frequency Readout Accuracy | Start, Stop, Center, Marker | ±(marker frequency indication X frequency reference accuracy + 10% x RBW + frequency resolution) |
|  | Trace points | Max 601 points, min 6 points |
| Marker Frequency Counter | Resolution | 1 Hz, 10 Hz, 100 Hz, 1 kHz |
|  | Accuracy | ±(marker frequency indication X frequency reference accuracy + counter resolution) | RBW/Span >=0.02 ; Mkr level to DNL>30 dB |
| Frequency Span | Range | 0 Hz (zero span), 100 Hz to 3 GHz |  |
|  | Resolution | 1 Hz |  |
|  | Accuracy | ± frequency resolution1 | RBW: Auto; |
| Phase Noise | Offset from Carrier |  | Fc =1 GHz; RBW = 1 kHz, VBW = 10 Hz; Average ≥ 40 |
|  | 10 kHz | <-88 dBc/Hz | Typical |
|  | 100 kHz | <-95 dBc/Hz | Typical |
|  | 1 MHz | <-113 dBc/Hz | Typical |
| Resolution Bandwidth (RBW) Filter | Filter Bandwidth | 1 Hz to 1 MHz in 1-3-10 sequence | -3dB bandwidth |
|  |  | 200 Hz, 9 kHz, 120 kHz, 1MHz | -6dB bandwidth |
|  | Accuracy | ± 8%, RBW = 1MHz | Nominal |
|  |  | ± 5%, RBW < 1MHz | Nominal |
|  | Shape Factor | < 4.5:1 | Normal Bandwidth ratio: -60dB:-3dB |
| Video Bandwidth (VBW) Filter | Filter Bandwidth | 1 Hz to 1 MHz in 1-3-10 sequence | -3dB bandwidth |
| Amplitude |
| Amplitude Range | Measurement Range | 100 kHz to 1 MHz | Displayed Average Noise Level (DANL) to 18 dBm |
|  |  | 1 MHz to 10 MHz | DANL to 21 dBm |
|  |  | 10 MHz to 3 GHz | DANL to 30 dBm |
| Attenuator | Input Attenuator Range | 0 to 50 dB, in 1 dB step | Auto or manual setup |
| Maximum Safe Input Level | Average Total Power | ≤ +33 dBm | Input attenuator ≥10 dB |
|  | DC Voltage | ± 50 V |  |
| 1 dB Gain Compression | Total Power at 1st Mixer | > 0 dBm | Typical;Fc ≥ 50 MHz; preamp. off |
|  | Total Power at the Preamp | > -22 dBm | Typical;Fc ≥ 50 MHz; preamp. on |
|  |  | mixer power level (dBm)= input power (dBm)-attenuation (dB) |
| Displayed Average Noise Level (DANL) | Preamp off | 0 dB attenuation; RF Input is terminated with a 50Ω load. RBW 10 Hz; VBW 10 Hz; span 500 Hz; reference level = -60dBm; trace average ≥ 40 |
|  | 9 kHz to 100 kHz | < -93 dBm | Nominal |
|  | 100 kHz to 1 MHz | < -90 dBm - 3 x (f/100 kHz) dB |  |
|  | 1 MHz to 2.7 GHz | < -122 dBm |  |
|  | 2.7 GHz to 3 GHz | < -116 dBm |  |
|  | Preamp on | 0 dB attenuation; RF Input is terminated with a 50Ω load ; RBW 10 Hz; VBW 10Hz; span 500 Hz; reference level = -60dBm; trace average ≥ 40 |
|  | 100 kHz to 1 MHz | < -108 dBm - 3 x (f/100 kHz) dB | Nominal |
|  | 1 MHz to 10 MHz | < -142 dBm |  |
|  | 10 MHz to 3 GHz | < -142 dBm + 3 x (f/1 GHz) dB |  |
| Level Display Range | Scales | Log, Linear |  |
|  | Units | dBm, dBmV, dBuV, V, W |  |
|  | Marker Level Readout | 0.01 dB | Log scale |
|  |  | 0.01 % of reference level | Linear scale |
|  | Level Display Modes | Trace, Topographic, Spectrogram | Single / split Windows |
|  | Number of Traces | 4 |  |
|  | Detector | Positive-peak, negative-peak, sample, normal, RMS(not Video) | Can be setup for each trace separately |
|  | Trace Functions | Clear & Write, Max/Min Hold, View, Blank, Average |  |
| Absolute Amplitude Accuracy | Absolute Point | Center=160 MHz ; RBW 10 kHz; VBW 1 kHz; span 100 kHz; log scale; 1 dB/div; peak detector; 23°C±1°C; Signal at Reference Level |
|  | Preamp off | ± 0.3 dB | Ref level 0 dBm; 10 dB RF attenuation |
|  | Preamp on | ± 0.4 dB | Ref level -30 dBm; 0 dB RF attenuation |
| Frequency Response | Preamp off | Attenuation: 10 dB; Reference: 160 MHz; 20 to 30°C |
|  | 100 kHz to 2.0 GHz | ± 0.5 dB |  |
|  | 2GHz to 3 GHz | ± 0.7 dB |  |
|  | Preamp on | Attenuation: 0 dB; Reference: 160 MHz; 20 to 30°C |
|  | 1 MHz to 2 GHz | ± 0.6 dB |  |
|  | 2 GHz to 3 GHz | ± 0.8 dB |  |
| Attenuation Switching Uncertainty | Attenuator setting | 0 to 50 dB in 1 dB step |  |
|  | Uncertainty | ± 0.25 dB | reference: 160 MHz, 10dB attenuation |
| RBW Filter Switching Uncertainty | 1 Hz to 1 MHz | ± 0.25 dB | reference : 10 kHz RBW |
| Level Measurement Uncertainty | Overall Amplitude Accuracy | ± 1.5 dB | 20 to 30°C; frequency > 1 MHz; Signal input 0 to -50 dBm; Reference level 0 to -50 dBm; Input attenuation 10 dB; RBW 1 kHz; VBW 1 kHz; after cal; Preamp Off |
|  |  | ± 0.5 dB | Typical |
| Spurious Response | Second Harmonic Intercept |  | Preamp off; signal input -30dBm; 0 dB attenuation |
|  |  | +35 dBm | Typical; 10 MHz < fc < 775 MHz |
|  |  | +60 dBm | Typical; 775 MHz ≤ fc < 1.625 GHz |
|  | Third-order Intercept |  | Preamp off; signal input -30dBm; 0 dB attenuation |
|  |  | > 1dBm | 300 MHz to 3 GHz |
|  | Input Related Spurious | < -60 dBc | Input signal level -30 dBm, Att. Mode, Att=0dB; 20-30ºC |
|  | Residual Response (inherent) | <-90 dBm | Input terminated; 0 dB attenuation; Preamp off |
| Sweep |
| Sweep Time | Range | 204 us to 1000 s | Span > 0 Hz |
|  |  | 50 us to 1000 s | Span = 0 Hz; Min Resolution = 10 us |
|  | Sweep Mode | Continuous; Single |  |
|  | Trigger Source | Free run; Video; External |  |
|  | Trigger Slope | Positive or negative edge |  |
| RF Preamplifier |
|  | Frequency Range | 1 MHz to 3 GHz |  |
|  | Gain | 18 dB | Nominal |
|  |  |  | (installed as standard) |
| Front Panel Input/Output |
| RF Input | Connector Type | N-type female |  |
|  | Impedance | 50 ohm | Nominal |
|  | VSWR | <1.6 :1 | 300 kHz to 3 GHz; Input attenuator ≥ 10 dB |
| Power for Option | Connector Type | SMB male |  |
|  | Voltage/Current | DC +7V / 500 mA max | With short-circuit protection |
| USB Host | Connector Type | A plug |  |
|  | Protocol | Version 2.0 | Supports Full/High/Low speed |
| MicroSD Socket | Protocol | SD 1.1 |  |
|  | Supported Cards | microSD, microSDHC | Up to 32GB capacity |
| Rear Panel Input/Output |
| Reference Output | Connector Type | BNC female |  |
|  | Output Frequency | 10 MHz | Nominal |
|  | Output Amplitude | 3.3V CMOS |  |
|  | Output Impedance | 50 ohm |  |
| Reference Input | Connector Type | BNC female |  |
|  | Input Reference Frequency | 10 MHz |  |
|  | Input Amplitude | -5 dBm to +10 dBm |  |
|  | Frequency Lock Range | Within ± 5 ppm of the input reference frequency |  |
| Alarm Output | Connector Type | BNC female | Open-collector |
| Trigger Input/ Gated Sweep Input | Connector Type | BNC female |  |
|  | Input Amplitude | 3.3V CMOS |  |
|  | Switch | Auto selection by function |  |
| LAN TCP/IP Interface | Connector Type | RJ-45 |  |
|  | Base | 10Base-T; 100Base-Tx; Auto-MDIX |
| USB Device | Connector Type | B plug | For remote control only; supports USB TMC |
|  | Protocol | Version 2.0 | Supports Full/High/Low speed |
| IF Output | Connector Type | SMA female |  |
|  | Impedance | 50 ohm | Nominal |
|  | IF Frequency | 886 MHz | Nominal |
|  | Output level | -25 dBm | 10 dB attenuation; RF input: 0 dBm @ 1 GHz |
| Earphone Output | Connector Type | 3.5mm stereo jack, wired for mono operation |
| Video Output | Connector Type | DVI-I ( integrated analog and digital) , Single Link. Compatible with VGA or HDMI standard through adapter |
| RS232 Interface | Connector Type | D-sub 9-pin female | Tx,Rx,RTS,CTS |
| GPIB Interface (Optional) | Connector Type | IEEE-488 bus connector |  |
| AC Power Input | Power Source | AC 100 V to 240 V, 50 / 60 Hz Auto range selection |
| Battery Pack (Optional) | Battery pack | 6 cells, Li-Ion rechargeable, 3S2P | With UN38.3 Certification |
|  | Voltage | DC 10.8 V |  |
|  | Capacity | 5200 mAh / 56Wh |  |
| General |
|  | Internal Data storage | 16 MB nominal |  |
|  | Power Consumption | <65 W |  |
|  | Warm-up Time | < 30 minutes |  |
|  | Temperature Range | +5 °C to +45 °C | Operating |
|  |  | -20 °C to + 70 °C | Storage |
|  | Weight | 4.5 kg (9.9 lb) | Inc. all options (Basic+TG+GPIB+Battery) |
|  | Dimensions | 210 x 350 x 100 (mm) | Approximately |
|  |  | 8.3 x 13.8 x 3.9 (in) |  |
| Tracking Generator (Optional) |
|  | Frequency Range | 100 kHz to 3 GHz |  |
|  | Output Power | -50 dBm to 0 dBm in 0.5 dB steps |
|  | Absolute Accuracy | ± 0.5 dB | @160 MHz, -10 dBm, Source attenuation 10 dB, 20 to 30°C |
|  | Output Flatness | Referenced to 160 MHz, -10 dBm |
|  |  | 100 kHz to 2 GHz | ± 1.5 dB |
|  |  | 2 GHz to 3 GHz | ± 2 dB |
|  | Output Level Switching Uncertainty | ± 0.8 dB | Referenced to -10 dBm |
|  | Harmonics | < -30 dBc | Typical, output level = -10 dBm |
|  | Reverse Power | +30 dBm max. |  |
|  | Connector type | N-type female |  |
|  | Impedance | 50 ohm | Nominal |
|  | Output VSWR | < 1.6:1 | 300 kHz to 3 GHz, source attenuation ≥ 12 dB |