PCS-1000I Specifications

The specifications apply when the PCS-1000I is powered on for at least 30 minutes under +20°C~+30°C.

| PCS-1000I Specifications |
| --- |
| General |
| Power Supply | 100 V / 120 V / 220 V / 240 V ±10% |
| Power Line Frequency | 50/60 Hz |
| Operating Environment | Full accuracy for0 ℃to55 ℃, Full accuracy to 80% R.H. at40 ℃ |
| Storage Environment | -40 ℃to70 ℃ |
| Power Consumption | Max 35VA |
| Dimensions | 210 mm(W) *80 mm(H) *390mm(D) |
| Weight | Approxiately5 kg |
| DC Characteristics |
| DC Voltage |
| Range | 1 Year23 ℃±5 ℃ | Temperature Coefficient/℃ |
| 200.0000 mV |  | 0.0050 + 0.0035 | 0.0005 + 0.0005 |
| 2.000000 V |  | 0.0050 + 0.0010 | 0.0005 + 0.0001 |
| 20.00000 V |  | 0.0050 + 0.0010 | 0.0005 + 0.0001 |
| 200.0000 V |  | 0.0050 + 0.0010 | 0.0005 + 0.0001 |
| 1000.000 V |  | 0.0050 + 0.0020 | 0.0005 + 0.0001 |
| Accuracy specification : ± (% of reading + % of range) |
| Voltage input Resistance:10MΩ for all DC Voltage ranges |
| DC Current |
| Range | Burden Voltage | 1 Year23 ℃±5 ℃ | Temperature Coefficient/℃ |
| 30.00000 mA | <0.4 V | 0.01 + 0.005 | 0.001 + 0.002 |
| 300.0000 mA | <0.5 V | 0.01 + 0.005 | 0.001 + 0.002 |
| 3.000000 A | <0.8 V | 0.01 + 0.005 | 0.001 + 0.002 |
| 30.00000 A*1 | <0.8 V | 0.01 + 0.005 | 0.001 + 0.002 |
| 300.0000 A*1 | <0.8 V | 0.02 + 0.005 | 0.001 + 0.002 |
| Accuracy specification : ± (% of reading + % of range) |
| Isolated DC Current Monitor Accuracy |
| Range | Resolution( 6 1/2) | DC Accuracy | Temperature Coefficient/℃ |
| 30.00000 mA | 0.00001mA | 0.1+0.05 | 0.001 |
| 300.0000 mA | 0.0001mA | 0.1+0.05 | 0.001 |
| 3.000000 A | 0.000001A | 0.1+0.05 | 0.001 |
| 30.00000 A*1 | 0.00001A | 0.1+0.05 | 0.001 |
| 300.0000 A*1 | 0.0001A | 0.2+0.05 | 0.001 |
| Accuracy specification : ± (% of output + % of full scale) |
| Monitor output voltage for the full scale current = 3V |
| AC Characteristics |
| True RMS AC Voltage |
| Range | Frequency | 1 Year23 ℃±5 ℃ | Temperature Coefficient/℃ |
| 200.0000 mV | 45Hz - 2kHz 2kHz - 10kHz 10kHz - 20kHz | 0.5 + 0.05 1.0 + 0.05 2.0 + 0.10 | 0.005 + 0.005 |
| 2.000000 V |  |  | 0.005 + 0.005 |
| 20.00000 V |  |  | 0.005 + 0.005 |
| 200.0000 V |  |  | 0.005 + 0.005 |
| 600.000 V |  |  | 0.005 + 0.005 |
| Accuracy specification : ± (% of reading + % of range) |
| True RMS AC Current |
| Range | Frequency | 1 Year23 ℃±5 ℃ | Temperature Coefficient/℃ |
| 30.00000 mA | 45Hz - 2kHz 2kHz - 10kHz | 0.5 + 0.05 1.0 + 0.05 | 0.03 + 0.006 |
| 300.0000 mA |  |  | 0.03 + 0.006 |
| 3.000000 A |  |  | 0.03 + 0.006 |
| 30.00000 A*1 | 45Hz - 400Hz | 0.5 + 0.05 | 0.03 + 0.006 |
| 300.0000 A*1 |  |  | 0.03 + 0.006 |
| Accuracy specification : ± (% of reading + % of range) |
| Isolated AC Current Monitor Accuracy |
| Range | Frequency | AC Accuracy | Max Input Limit |
| 30.00000 mA | 45Hz - 2kHz 2kHz - 10kHz | 0.2 + 0.05 0.5 + 0.05 | 30 mA |
| 300.0000 mA |  |  | 300 mA |
| 3.000000 A |  |  | 3 A |
| 30.00000 A*1 | 45Hz - 400Hz | 0.5 + 0.05 | 30 A |
| 300.0000 A*1 |  |  | 300 A |
| Accuracy specification : ± (% of output + % of full scale) |
| Monitor output voltage for the full scale current = 3V |
| The specifications are only applicable when the input is 10% or greater of the full scale range |

*1 The accuracy for 30A/300A levels must be increased by a power factor of 8 ppm/watt.